## 0308
2022-08-03 | Agr
GA | 135 |  | Source | B | fps | 15 | gain | 20
folder | couch height | kV | mAs | focal spot
0005 | 8 | 100 | 50 | small
0006 | 5
0007 | 0
0008 | -5
0009 | -10
0010 | 8 |  |  |  | with ruler T, staples G, sharpener A
0011 | 8 |  |  |  | moved objects closer to centre
11/08/2022
Extended BOS Board
Original Couch Pos:
Height | -33
Long | 30
Lat | -28
Gantry Angle:
GA | 135 |  | Source | B
kV Settings:
kV | 100
mAs | 50
focal spot | small
source | B
Beamworks Settings:
fps | 15
gain | 20
script | CapturePeakMultiSpot4000.scr
Folder | Height
-25
-28
-33
-38
-43

## 1108
11/08/2022 |  |  |  |  | Mean Pixel Coordinates | [7 empty] | px
|  | Folder | Couch Height (cm) | Distance From Iso (mm) | A Edge | B Edge | G Edge | T Edge | A 10cm | B 10cm | G 10cm | T 10cm | AB Diff | GT Diff | AB 10cm | GT 10cm | AB Laser Centre | GT Laser Centre | AB mm/px | GT mm/px | A Edge (mm) | B Edge (mm) | G Edge (mm) | T Edge (mm) | AB kV Centre (mm) | GT kV Centre (mm)
Couch Accessories: |  | 0001 | -22 | 80 | 200 | 1352 | 196 | 1084 | 437 | 1165 | 236 | 960 | 1152 | 888 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -165.1099 | 151.3736 | -111.0497 | 134.2541 | -6.8681 | 11.6022
insert | Extended BOS Board | 0002 | -25 | 50 | 192 | 1353 | 204 | 1090 | 437 | 1165 | 236 | 960 | 1161 | 886 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -167.3077 | 151.6484 | -108.8398 | 135.9116 | -7.8297 | 13.5359
|  | 0003 | -30 | 0 | 186 | 1359 | 189 | 1092 | 437 | 1165 | 236 | 960 | 1173 | 903 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -168.956 | 153.2967 | -112.9834 | 136.4641 | -7.8297 | 11.7403
Original Couch Pos: |  | 0004 | -35 | -50 | 174 | 1371 | 179 | 1097 | 437 | 1165 | 236 | 960 | 1197 | 918 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -172.2527 | 156.5934 | -115.7459 | 137.8453 | -7.8297 | 11.0497
Height | -30 | 0005 | -15 | 150 | 217 | 1331 | 209 | 1064 | 437 | 1165 | 236 | 960 | 1114 | 855 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -160.4396 | 145.6044 | -107.4586 | 128.7293 | -7.4176 | 10.6354
Long | 30 | 0006 | -10 | 200 | 234 | 1319 | 218 | 1060 | 437 | 1165 | 236 | 960 | 1085 | 842 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -155.7692 | 142.3077 | -104.9724 | 127.6243 | -6.7308 | 11.326
Lat | -28 | 0007 | -5 | 250 | 248 | 1310 | 224 | 1048 | 437 | 1165 | 236 | 960 | 1062 | 824 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -151.9231 | 139.8352 | -103.3149 | 124.3094 | -6.044 | 10.4972
|  | 0008 | 0 | 300 | 261 | 1291 | 232 | 1035 | 437 | 1165 | 236 | 960 | 1030 | 803 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -148.3516 | 134.6154 | -101.105 | 120.7182 | -6.8681 | 9.8066
Gantry Angle: |  | 0009 | 5 | 350 | 268 | 1286 | 242 | 1031 | 437 | 1165 | 236 | 960 | 1018 | 789 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -146.4286 | 133.2418 | -98.3425 | 119.6133 | -6.5934 | 10.6354
GA | 135 | 0010 | 8 | 380 | 272 | 1284 | 244 | 1025 | 437 | 1165 | 236 | 960 | 1012 | 781 | 728 | 724 | 801 | 598 | 0.27473 | 0.27624 | -145.3297 | 132.6923 | -97.7901 | 117.9558 | -6.3187 | 10.0829
kV Settings: |  |  |  |  | A1000 | -117.978 | 224.695
kV | 100 | px1 | 209 |  | B1000 | 106.717
mAs | 50 | px2 | 279 |  | G1000 | -76.845 | 120.375
focal spot | small | pxhalf | 244 |  | T1000 | 43.53
source | B |  |  |  | 0 | 250
SID (mm) | 3700 |  |  |  | 0 | -250
SAD (mm) | 2700 |  |  |  | 0 | 0
AID (mm) | 1000
AB Field of View (mm) | 400
GT field of View (mm) | ???
Beamworks Settings:
fps | 15
gain | 20
script | CapturePeakMultiSpot4000.scr
G
T
B